## Slide 1: Introduction to X-ray Imaging James Finch

## Slide 2: Structure
What are X-rays?
Generating X-rays
Casting a shadow
Detecting X-rays
Computed Tomography (CT)

## Slide 3
What are X-rays?

## Slide 4: What are X-rays?
X-rays are electromagnetic radiation, like visible light and radio waves
They travel in straight lines
Once the source is switched off, they dissipate immediately
The energy of individual X-rays is measured in eV (electron-volts)
Nikon X-ray sources produce X-ray energies in the hundreds of thousands of eV: 30-450 keV
Image courtesy of Wikimedia Foundation

## Slide 5
Generating X-rays: X-ray Sources

## Slide 6: Generating X-rays: Reflection Target
Filament
Anode
Magnetic Lens
Target
X-rays
Output window
X/Y Coils
A hot filament emits electrons
High voltage between the filament and anode attracts and accelerates the electrons into a beam
X/Y coils steer the beam into the magnetic lens, which focuses the beam down to a few µm
The electrons impact the target, producing X-rays
  Over 99% of the electron beam’s power is transformed into heat.
  The target is liquid cooled.
The X-ray source is continuously pumped to a high vacuum. The output window seals the vacuum whilst allowing X-rays out

## Slide 7: Generating X-rays: Transmission Target
Filament
Anode
Magnetic Lens
Target Window
X-rays
X/Y Coils
The transmission target consists of a thin layer of tungsten on the inside surface of the output window; the target is the output window.
This allows:
A smaller focal spot size than reflection targets
Higher magnifications, as the focal spot is 0.25 mm behind the output window

## Slide 8: Generating X-rays: Rotating Reflection Target
The rotating reflection target focuses the electron beam onto the edge of a spinning disc.
This allows for much better heat dissipation.
Compared to the static reflection target, the rotating reflection target allows:
Twice the maximum power (450W)
Significantly smaller focal spot sizes at higher powers
The rotating target is operator-interchangeable with other targets.

## Slide 9: Generating X-rays: 320 Module
The 320 module applies up to +100kV to the target, boosting the maximum beam energy to 320kV
This allows:
Better penetration through dense samples
Higher maximum power of 320W
The 320 module is operator-interchangeable with other targets
Cathode
-220kV
Target
+100kV
320kV total

## Slide 10: Generating X-rays: 450kV Reflection Target
The 450kV reflection target applies up to +225kV to the target, bosting the maximum beam energy to 450kV.
This allows for better penetration through dense samples.
Cathode
-225kV
Target
+225kV
450kV total

## Slide 11: Generating X-rays: 450kV Rotating Target
The 450kV reflection target applies up to +225kV to the target, bosting the maximum beam energy to 450kV.
This allows for better penetration through dense samples.
The rotating target focuses the electron beam onto the edge of a spinning disc.
Compared to the static target, the rotating target allows significantly smaller focal spot sizes at higher powers.
Cathode
-225kV
Target
+225kV
450kV total

## Slide 12
Generating X-rays: Production Mechanisms

## Slide 13: Generating X-Rays: Bremsstrahlung Radiation
A high speed electron is deflected by the atoms in the target and loses energy
This energy is released as an X-ray
As electrons may be deflected by different amounts, Bremsstrahlung releases a continuous spectrum of X-ray energies up to the electron beam energy
Target + Atom
X-ray emission
e-
e-

## Slide 14: Generating X-rays: Characteristic Emission
e-
e-
e-
e-
Target
Nucleus
A high speed electron collides with an electron in the target atom and knocks it out of its orbital. This leaves a hole behind

## Slide 15: Generating X-rays: Characteristic Emission
Target
Nucleus
e-
e-
e-
e-
X-ray emission
A high speed electron collides with an electron in the target atom and knocks it out of its orbital. This leaves a hole behind
An electron in a higher-energy shell “falls down” to fill the lower-energy shell. This transition releases energy as an X-ray
As there are only specific energy levels that the orbital electrons can be in, there are only specific transitions possible – these are characteristic of the target material

## Slide 16: Generating X-rays: The X-ray Spectrum
X-ray energy / keV
Relative intensity
Accelerating voltage
controls maximum X-ray energy
Continuous Bremsstrahlung emission
Lowest energy X-rays don’t make it out of the source
Characteristic emission lines
Bremsstrahlung and Characteristic Emission both occur when electrons hit the target

## Slide 17: Multi-metal reflection target
Tungsten reflection target
Multi-metal reflection target
The standard reflection target is made from Tungsten (W) because it has intense characteristic emission around ~60keV.
At accelerating voltages below that, the characteristic emission no longer occurs.
The multi-metal reflection target offers a choice of 4 metals. The characteristic emission peaks of these are:
Tungsten (W): 58-69 keV
Silver (Ag): 22-26 keV
Molybdenum (Mo): 18-20 keV
Copper (Cu): 8-9 keV
Mo is the most commonly used alternative to W.

## Slide 18
Casting A Shadow

## Slide 19: Casting A Shadow
Just like light, X-rays travel in straight lines from the source. Where X-rays are absorbed by the sample, they leave a shadow image.
Unlike light, some X-rays will always penetrate the sample. Denser and thicker parts of the sample absorb more X-rays, so the darkness of the shadow tells us how much material is in the path of the X-rays.

## Slide 20: Casting A Shadow: Attenuation
Chart: Distance Penetrated (mm)
| 0 | 0.1 | 0.2 | 0.30000000000000004 | 0.4 | 0.5 | 0.6 | 0.7 | 0.79999999999999993 | 0.89999999999999991 | 0.99999999999999989 | 1.0999999999999999 | 1.2 | 1.3 | 1.4000000000000001 | 1.5000000000000002 | 1.6000000000000003 | 1.7000000000000004 | 1.8000000000000005 | 1.9000000000000006 | 2.0000000000000004 | 2.1000000000000005 | 2.2000000000000006 | 2.3000000000000007 | 2.4000000000000008 | 2.5000000000000009 | 2.600000000000001 | 2.7000000000000011 | 2.8000000000000012 | 2.9000000000000012 | 3.0000000000000013 | 3.1000000000000014 | 3.2000000000000015 | 3.3000000000000016 | 3.4000000000000017 | 3.5000000000000018 | 3.6000000000000019 | 3.700000000000002 | 3.800000000000002 | 3.9000000000000021 | 4.0000000000000018 | 4.1000000000000014 | 4.2000000000000011 | 4.3000000000000007 | 4.4000000000000004 | 4.5 | 4.5999999999999996 | 4.6999999999999993 | 4.7999999999999989 | 4.8999999999999986 | 4.9999999999999982
I=Io exp(-µt) | 1 | 0.90484 | 0.81873 | 0.74082 | 0.67032 | 0.60653 | 0.54881 | 0.49659 | 0.44933 | 0.40657 | 0.36788 | 0.33287 | 0.30119 | 0.27253 | 0.2466 | 0.22313 | 0.2019 | 0.18268 | 0.1653 | 0.14957 | 0.13534 | 0.12246 | 0.1108 | 0.10026 | 0.090718 | 0.082085 | 0.074274 | 0.067206 | 0.06081 | 0.055023 | 0.049787 | 0.045049 | 0.040762 | 0.036883 | 0.033373 | 0.030197 | 0.027324 | 0.024724 | 0.022371 | 0.020242 | 0.018316 | 0.016573 | 0.014996 | 0.013569 | 0.012277 | 0.011109 | 0.010052 | 0.0090953 | 0.0082297 | 0.0074466 | 0.0067379
I=Io exp(-µ't) | 1 | 0.81873 | 0.67032 | 0.54881 | 0.44933 | 0.36788 | 0.30119 | 0.2466 | 0.2019 | 0.1653 | 0.13534 | 0.1108 | 0.090718 | 0.074274 | 0.06081 | 0.049787 | 0.040762 | 0.033373 | 0.027324 | 0.022371 | 0.018316 | 0.014996 | 0.012277 | 0.010052 | 0.0082297 | 0.0067379 | 0.0055166 | 0.0045166 | 0.0036979 | 0.0030276 | 0.0024788 | 0.0020294 | 0.0016616 | 0.0013604 | 0.0011138 | 0.00091188 | 0.00074659 | 0.00061125 | 0.00050045 | 0.00040973 | 0.00033546 | 0.00027465 | 0.00022487 | 0.00018411 | 0.00015073 | 0.00012341 | 0.00010104 | 8.2724e-05 | 6.7729e-05 | 5.5452e-05 | 4.54e-05
I=Io exp(-µ''t) | 1 | 0.95123 | 0.90484 | 0.86071 | 0.81873 | 0.7788 | 0.74082 | 0.70469 | 0.67032 | 0.63763 | 0.60653 | 0.57695 | 0.54881 | 0.52205 | 0.49659 | 0.47237 | 0.44933 | 0.42741 | 0.40657 | 0.38674 | 0.36788 | 0.34994 | 0.33287 | 0.31664 | 0.30119 | 0.2865 | 0.27253 | 0.25924 | 0.2466 | 0.23457 | 0.22313 | 0.21225 | 0.2019 | 0.19205 | 0.18268 | 0.17377 | 0.1653 | 0.15724 | 0.14957 | 0.14227 | 0.13534 | 0.12873 | 0.12246 | 0.11648 | 0.1108 | 0.1054 | 0.10026 | 0.095369 | 0.090718 | 0.086294 | 0.082085
X-rays are absorbed as an exponential decay.
I.e., if one millimetre of a material absorbs half of the X-rays, then the next millimetre will absorb half of what remains: one quarter of the original beam.
The density of electrons in the sample material affects how quickly X-rays are absorbed; samples with a higher atomic (Z) number or higher density will absorb more X-rays.

## Slide 21: Casting A Shadow: Attenuation
Penetration
X-ray energy
Higher energy X-rays penetrate further
By increasing the beam energy, we get higher energy X-rays which are more penetrating.
By reducing the beam energy, we get only lower energy X-rays which are absorbed easier and give better contrast.
We can also adjust the beam current, which gives us more or fewer X-rays of the same energy. This affects brightness but not penetration.

## Slide 22: Casting A Shadow: Effect of changing kV
Relative intensity
200
The energy of each individual X-ray is measured in keV (kilo electron-volts).
A 200kV accelerating voltage produces X-rays with energies up to but never more than 200keV. The beam is described as 200kVp (peak).
Most X-ray emission is at lower energies.
Although increasing the accelerating voltage (kV) does produce more high-energy X-rays, it doesn’t affect penetration as much as you might expect.
200kVp
61keV mean
66keV mean
225kVp
225
X-ray energy / keV

## Slide 23: Casting A Shadow: Effect of filters
X-ray energy / keV
Relative intensity
Adding a metal filter in front of the X-ray source will remove low-energy X-rays from the beam, increasing the mean energy and therefore penetration.
Because filters remove X-rays, they also make the beam significantly darker (increasing noise).
200kVp
no filter
200kVp
2mm Copper
65keV mean
112keV mean
200

## Slide 24: Casting A Shadow: Geometric Magnification
The X-ray beam starts as a point, and spreads out as a cone to the detector.
X-rays are extremely hard to bend, so we don’t use a lens like with visible light. Instead, we move the sample within the cone beam to take advantage of geometric magnification.
A sample placed close to the detector will not be magnified much
A sample placed close to the X-ray source will be highly magnified
You can test this with the light on your phone and your hand!
Low magnification: Sample far from source
High magnification: Sample close to source

## Slide 25: Casting A Shadow: Geometric Unsharpness
Small focal spot
Large focal spot
blurring
At high geometric magnifications, any blurring (unsharpness) caused by the focal spot is magnified.
It is essential to have a small focal spot size to avoid blurring at high magnifications.
Microfocus X-ray sources have a focal spot in the micron-range, providing sharp images at extremely high geometric magnifications.

## Slide 26
Detecting X-rays

## Slide 27: Detecting X-rays
Most modern X-ray systems use a Digital Flat Panel Detector
These consist of two layers:
X-rays are absorbed by the scintillator crystal and re-emitted as visible light
The visible light is detected by photodiodes
A digital flat panel is effectively a very large digital camera sensor, with a scintillator in front.

## Slide 28: Detecting X-rays: Effective Pixel Size
2x geometric magnification
100µm
50µm
If the sample is halfway between the source and detector, the image will be twice as large as the sample; there is 2x geometric magnification.
This means that a 100µm pixel represents 50µm in the sample. The effective pixel size is 50µm.
To reduce blurring caused by the X-ray source, we should try to keep the X-ray focal spot size less than the effective pixel size.

## Slide 29: Detecting X-rays: Noise
1 second
16 seconds
128 seconds
All images contain some amount of noise. This is also true of X-ray images.
To improve the Signal-to-Noise Ratio, you can use a brighter X-ray source or a longer exposure time; either of these will increase the X-ray flux – the number of X-rays reaching the detector.
The SNR improves with the square root of flux, so an image that takes 16x as long to acquire will be 4x less noisy. The more noise you remove, the harder it gets to remove more noise.

## Slide 30
Computed Tomography

## Slide 31: What is CT?
Computed Tomography (CT) provides a 3D map of the X-ray absorption throughout the sample.
By rotating the sample 360 degrees whilst acquiring images, we get projection images from every angle.
These projection images are then reconstructed into a 3D volume.

## Slide 32: CT: Acquisition, Reconstruction and Analysis
Acquisition
Projection Images
Reconstruction
2D Slices + 3D Rendering
Analysis
Results

## Slide 33: CT: Number of projection images
500 projections
1000 projections
2000 projections
3000 projections
During the scan, no part of the sample should move more than one pixel across the detector between projection images.
If it does, fine detail will be lost to fringes. This affects the edges of the scan the most.
The optimum number of projections depends on the number of pixels horizontally on the detector that the sample rotates across.
Acquiring more projections than necessary will not improve feature recognition further, however the extra data will reduce noise.

## Slide 34: X.Tend Helical Acquisition
With X.Tend, the stage rotates and Y axis moves up at the same time. This is ideal for long parts with a small diameter.
Reduces artefacts
Provides higher resolution
Allows taller samples
Circular CT
X.Tend

## Slide 35: Limited Angle CT
Limited Angle CT allows more flexibility to scan regions of large samples by avoiding collision with the source or detector
Higher magnification for better resolution at edges of large samples
Lower magnification for larger field-of-view at edges of large sample

## Slide 36
Beam Hardening

## Slide 37: Beam Hardening
High-energy (hard) X-rays are more likely to penetrate deeper through the sample than low-energy (soft) X-rays.
As the X-rays travel through the sample, the lower energy X-rays are removed from the beam leaving behind only higher energy X-rays – the beam gets harder.
This means that we get much more X-ray absorption at the edges of the sample. This is called beam hardening.

## Slide 38: Beam Hardening
The bright-to-dark gradient is a common symptom of beam hardening.
This can be seen as “cupping” in a line profile.

## Slide 39: Beam Hardening
Disk of 40mm diameter (orange) and considering a chord of 4.5mm (blue)

## Slide 40: Beam Hardening
Left to right: 0, 1, 2.5mm Cu

## Slide 41: Beam Hardening – hardware correction
Left to right: 0, 1, 2.5mm Cu

## Slide 42: Beam Hardening – software correction
Left to right: 0mm Cu, 0mm Cu with beam hardening correction
Greyvalue difference: 54.7% uncorrected, 7.9% corrected